## Cap. 3
Riferimento | Dettagli | Subcategory | Note | Controlli | Controlli_EN
3 Secure Design and Manufacture
3.3. Security Capabilities | Where there is an impact on safety or effectiveness12, manufacturers shall select the most appropriate risk control solution, in the following order of priority: a) Eliminate or reduce risks as far as possible through safe design and manufacture; b) Where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; c) Provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. | ID.RM-1 |  | ID.RM-1-01: i fabbricanti selezionano la soluzione di controllo del rischio più appropriata, nel seguente ordine di priorità: a) Eliminano o riducono i rischi attraverso una progettazione e una fabbricazione sicure; b)adottano adeguate misure di protezione, anche allarmi se necessario, in relazione a rischi che non possono essere eliminati; c) Forniscono informazioni per la sicurezza (avvertenze/precauzioni/controindicazioni) e formazione agli utenti. | ID.RM-1-01: Manufacturers shall select the most appropriate risk control solution, in the following order of priority: (a) Eliminate or reduce risks through safe design and manufacture; (b)take appropriate protective measures, including alerts where necessary, in relation to risks that cannot be eliminated; c) Provide safety information (warnings/precautions/contraindications) and training to users.
3.5. Security Benefit Risk Analysis | Risk acceptance criteria should be established by the manufacturer and documented to guide the appropriate measures for mitigating security risks. Those criteria relate to the intended purpose and operational environment. | ID.RM-3 |  | ID.RM-3-01: I criteri di accettazione del rischio sono stabiliti dal fabbricante e documentati per mitigare i rischi per la sicurezza. | ID.RM-3-01: Risk acceptance criteria are established by the manufacturer and documented to mitigate safety risks.
3.6. Minimum IT Requirements | It is the manufacturers’ responsibility to determine the minimum requirements for the operating environment as regards IT network characteristics and IT security measures that could not be implemented through the product design. | N/A
| IT security measures may refer to any applicable technical and/or organisational measures for managing IT security risks related to the operating environment. | PR.IP-1 |  | PR.IP-1-01: Le misure di sicurezza informatica si riferiscono a qualsiasi misura tecnica e/o organizzativa applicabile per la gestione dei rischi per la sicurezza informatica relativi all'ambiente operativo. | PR. IP-1-01: Cybersecurity measures refer to any technical and/or organizational measures applicable to the management of cybersecurity risks related to the operating environment.
| The manufacturer shall provide clear documentation of the device’s instructions for use, including IT security features/configurations (if applicable), and clear instructions for the IT security controls related to the operating environment, including product specifications, compatibilities, recommended IT security measures, IT environment configuration (e.g. traffic control), etc. | N/A
| Any minimum requirements concerning hardware, IT networks characteristics and IT security measures for the operating environment should be defined on the basis of the following principles: - Any proposed IT security requirement for the operating environment should be based on the risk assessment conducted for the medical device. | ID.GV-1 | da rivedere | ID.GV-1-01: Qualsiasi requisito di sicurezza IT proposto per l'ambiente operativo è basato sulla valutazione del rischio condotta per il dispositivo medico. | ID.GV-1-01: Any proposed IT security requirements for the operating environment are based on the risk assessment conducted for the medical device.
| - The medical device should be as autonomous as possible in terms of IT security and sole reliance on the existence of any IT security requirements on the operating environment should be kept to a minimum and reflect the manufacturer’s assumptions on the baseline environment security for the secure operation of the medical device. | N/A
| - The manufacturer’s assumptions regarding the IT security of the operating environment shall be clearly documented in the instructions for use and may refer to best practice security standards. | PR.IP-1 |  | PR.IP-1-02: Le ipotesi del produttore in merito alla sicurezza informatica dell'ambiente operativo sono chiaramente documentate nelle istruzioni per l'uso e fanno riferimento a standard di sicurezza delle migliori pratiche. | PR. IP-1-02: The manufacturer's assumptions regarding the cybersecurity of the operating environment are clearly documented in the instructions for use and refer to best practice security standards.
| - In accordance with the principle of layered security, IT security measures foreseen for the operating environment in general should not serve the purpose of compensating security controls for medical device vulnerabilities, unless there is sufficient justification. In cases where the medical device relies on the operating environment to provide important IT security controls, this should be stated in the accompanying technical documentation. | N/A
| The operator must be in line with national and EU regulations (e.g. GDPR). | N/A
| The operating environment must provide physical security for the medical device via security measures such as: - Regulated and authenticated physical access enforced via suitable technical measures (e.g. badges) | PR.AC-2 |  | PR.AC-2-01: l'accesso fisico è regolamentato e autenticato mediante idonee misure tecniche (es. badge) | PR. AC-2-01: physical access is regulated and authenticated through appropriate technical measures (e.g. badges)
| - Physical security policy defining roles and access rights, including for physical access to the medical device | PR.AC-2; PR.IP-5 |  | PR.AC-2-02: la politica di sicurezza fisica definisce ruoli e diritti di accesso anche per l'accesso fisico al dispositivo medico | PR. AC-2-02: Physical security policy defines roles and access rights also for physical access to the medical device
| - Use of segregated, secure areas with appropriate access controls | PR.AC-5 |  | PR.AC-5-01: sono utilizzate aree segregate e sicure con adeguati controlli di accesso | PR. AC-5-01: segregated and secure areas with adequate access controls are used
| The operating environment must include appropriate security controls such as: - User access management (credentials for accessing software applications or devices, user access policy, etc.) | PR.AC-1 |  | PR.AC-1-01: L'ambiente operativo include la gestione degli accessi utenti (credenziali di accesso ad applicazioni o dispositivi software, policy di accesso utenti, ecc.) | PR. AC-1-01: The operating environment includes user access management (access credentials to applications or software devices, user access policies, etc.)
| - Antivirus / anti-malware software - Application whitelisting / system hardening | DE.CM-4 | era gia stato mappato in SW sicurezza di base | DE.CM-4-01: L'ambiente operativo include la gestione di: - Software antivirus/antimalware - Whitelist dell'applicazione / rafforzamento del sistema | DE. CM-4-01: The operating environment includes the management of: - Anti-virus/antimalware software - Application whitelist / system hardening
| - Firewall | PR.AC-5 |  | PR.AC-5-02: L'ambiente operativo include la gestione di Firewall | PR.AC-5-02: L'ambiente operativo include la gestione di Firewall
| - Exclusive use of genuine software and ban of all illegitimate software and applications | PR.DS-6 |  | PR.DS-6-01: L'ambiente operativo permette l'uso esclusivo di software originale e vieta tutti i software e le applicazioni illegittimi o Misure di gestione delle sessioni (es. timeout delle sessioni)" | PR.DS-6-01: L'ambiente operativo permette l'uso esclusivo di software originale e vieta tutti i software e le applicazioni illegittimi o Misure di gestione delle sessioni (es. timeout delle sessioni)"
| - Session management measures (e.g. session timeouts) | PR.AC-3 |  | PR.AC-3-01: L'ambiente operativo implementa misure di gestione delle sessioni (es. timeout delle sessioni) | PR.AC-3-01: L'ambiente operativo implementa misure di gestione delle sessioni (es. timeout delle sessioni)
| The operating environment must provide control and security of network traffic via appropriate measures, such as: - Network segmentation - Traffic filtering - Data encryption | PR.AC-5 |  | PR.AC-5-03: L'ambiente operativo fornisce il controllo e la sicurezza del traffico di rete attraverso Segmentazione della rete 'PR.AC-5-04: L'ambiente operativo fornisce il controllo e la sicurezza del traffico di rete attraverso Filtraggio del traffico 'PR.AC-5-05: L'ambiente operativo fornisce il controllo e la sicurezza del traffico di rete attraversoCrittografia dei dati | PR.AC-5-03: L'ambiente operativo fornisce il controllo e la sicurezza del traffico di rete attraverso Segmentazione della rete 'PR.AC-5-04: L'ambiente operativo fornisce il controllo e la sicurezza del traffico di rete attraverso Filtraggio del traffico 'PR.AC-5-05: L'ambiente operativo fornisce il controllo e la sicurezza del traffico di rete attraversoCrittografia dei dati
| Specifically for the workstations connected to the medical device, appropriate security measures may include: - Operating system hardening and application whitelisting | PR.DS-1, PR-IP-1 |  | PR.DS-1-01: le misure di sicurezza comprendono il rafforzamento del sistema operativo e whitelist delle applicazioni | PR. DS-1-01: Security measures include OS hardening and application whitelisting
| - Memory protection measures to block arbitrary code execution | PR.DS-1, PR-IP-1 |  | PR.DS-1-02: le misure di sicurezza comprendono misure di protezione della memoria per bloccare l'esecuzione di codice arbitrario | PR. DS-1-02: Security measures include memory protection measures to block arbitrary code execution
| - Compatibility of medical device management software with security solutions that counter malicious code | N/A
| - use of strong passwords | PR.AC-1: |  | PR.AC-1-02: le misure di sicurezza comprendono l'uso di password complesse | PR. AC-1-02: Security measures include the use of strong passwords
| - Install only software programmes necessary for the intended use of the operating environment. | N/A
| For cases when the operating environment is a complex system integrating multiple medical devices and other systems, appropriate measures to limit the propagation of an attack may include: - Partitioning mechanisms and network / traffic segmentation | PR.AC-5 |  | PR.AC-5-04: In un sistema complesso (che integra più dispositivi medici e altri sistemi), le misure per limitare la propagazione di un attacco includono meccanismi di partizionamento e segmentazione rete/traffico | PR. AC-5-04: In a complex system (integrating multiple medical devices and other systems), measures to limit the propagation of an attack include partitioning mechanisms and network/traffic segmentation
| - Software integrity checks and device authentication mechanisms | PR.DS-6 |  | PR.DS-6-02: 'In un sistema complesso (che integra più dispositivi medici e altri sistemi), le misure per limitare la propagazione di un attacco includono controlli di integrità del software e meccanismi di autenticazione del dispositivo | PR. DS-6-02: 'In a complex system (integrating multiple medical devices and other systems), measures to limit the propagation of an attack include software integrity checks and device authentication mechanisms.
| To ensure that the security posture of the operating environment and of the device itself remain at a suitable level, appropriate provisions regarding patch management should be in place, such as: - The operating environment should support patching without compromising interoperability/compatibility - The operator should have appropriate patch management processes to ensure that security patches for medical devices are deployed in a timely manner - The operator should have appropriate patch management processes to ensure that the operating environment (e.g. operating systems, applications) is up-to-date in terms of security | PR.IP-7 |  | PR.IP-7-01: L'ambiente operativo supporta l'applicazione di patch senza compromettere l'interoperabilità/compatibilità PR.IP-7-02: L'operatore dispone di processi di gestione delle patch appropriati per garantire che le patch di sicurezza per i dispositivi medici siano implementate in modo tempestivo PR.IP-7-03: L'operatore dispone di processi di gestione delle patch per garantire che l'ambiente operativo (ad es. sistemi operativi, applicazioni) sia aggiornato in termini di sicurezza
| Elements of the operating environment interacting with (e.g. other devices) or required for the operation of medical devices (e.g. OS) should ensure interoperability and shall not impair the specified performance of the medical device | N/A
3.8. Lifecycle Aspects | Addressing cybersecurity risks at the design stage can help mitigate cybersecurity risks that could contribute to a breach in the confidentiality, a compromise in the integrity and availability of the medical device and its data, or intentional unauthorised access to the medical device and/or the network. | N/A
| During the support lifetime of the device, the manufacturer should put in place a process to gather post-market information with respect to the security of the device (see also Chapter 6). This process should take into account: 1. Security incidents directly related to medical device software 2. Security Vulnerabilities that are related to the medical device hardware/software and the 3rd party hardware/software used with the medical device. 3. Changes in the threat landscape, including interoperability aspects | PR.IP-2 |  | PR.IP-2-01: il produttore mette in atto un processo per raccogliere informazioni post-commercializzazione in merito alla sicurezza del dispositivo | PR. IP-2-01: The manufacturer puts in place a process to collect post-market information regarding the safety of the device

## Cap. 5
Riferimento | Dettagli | Subcategory | Note | Controlli | Controlli_EN
3 Secure Design and Manufacture
5.1. Post-market surveillance system | Handling and remediation of cybersecurity incidents and vulnerabilities reported through the post- market surveillance and vigilance systems shall be carried out conforming to the methodologies described in Chapter 3.2 of this guidance, with regards to: - Assess the need for reporting serious and non-serious incidents and of carrying-out field safety corrective actions; - Update the Verification and Validation; - Update the Technical Documentation. - Enhancing security capabilities; - Update the original Security Risk Assessment; - Update the original Security Benefit Risk Analysis" | N/A
5.2. Vigilance | The manufacturer is responsible for reporting all serious incidents and field safety corrective actions (FSCA) to the CA | ID.AM-6 |  | ID.AM-6-01: Il produttore è responsabile della segnalazione di tutti gli incidenti gravi e delle azioni correttive di sicurezza sul campo alla CA | ID.AM-6-01: The manufacturer is responsible for reporting all serious incidents and field safety corrective actions to the CA
| Manufacturers are obliged to conduct investigations as soon as they are informed that a serious incident has taken place. As such, a risk assessment of the incident is conducted and if needed, a FSCA will be implemented in order to minimize the risk of the device. | ID.RA-5 |  | ID.RA-5-01: I produttori conducono indagini in seguito al verifivarsi di un incidente grave. ID.RA-5-02: I produttori conducono una valutazione del rischio dell'incidente e, se necessario, implementano una FSCA (Field Safety Corrective actions) per ridurre al minimo il rischio del dispositivo. | ID.RA-5-01: Manufacturers conduct investigations following the occurrence of a serious incident. ID.RA-5-02: Manufacturers conduct an accident risk assessment and, if necessary, implement a Field Safety Corrective actions (FSCA) to minimize the risk of the device.
| The manufacturers shall carry out investigations of serious incidents related to a cybersecurity incident in order to provide a comprehensive description of the serious incident, including: a) a description of the serious incident including any relevant information that might impact the understanding or evaluation of the serious incident, i.e. information is compromised or information is threatened; b) a description of the health effects (if applicable), i.e. clinical signs, symptoms, conditions as well as the overall health impact. | N/A
| Incidents that have cybersecurity related incident root causes are subject to Trend Reporting under the Medical Devices Regulations | RS.CO-3 |  | RS.CO-3-01: Gli incidenti che hanno cause profonde legate alla sicurezza informatica sono soggetti a Trend Reporting | RS. CO-3-01: Incidents that have root causes related to cybersecurity are subject to Trend Reporting

## Annex I.I
Riferimento | Dettagli | Subcategory | Note | Controlli | Controlli_EN
7. Annex I – Mapping of IT security requirements to NIS Directive Cooperation Group measures | Conduct a risk assessment and impact assessment. Introduction of medical devices in the environment should be subject to such a risk assessment. | ID.RA-5 |  | ID.RA-5-03: è condotta una valutazione del rischio e una valutazione dell'impatto in seguito all'introduzione di dispositivi medici nell'ambiente. | ID.RA-5-03: A risk assessment and an impact assessment following the introduction of medical devices into the environment shall be conducted.
| A set of baseline IT security policies should be defined, approved by management and communicated to employees and relevant external parties. Examples of such policies include: - Define clear roles for users of critical systems and medical devices, segregate duties o Define information classification levels and label devices handling such information accordingly - Acceptable use policy - Password policy - Change Management policy - Establish a backup policy on selected critical systems. - Establish a vendor/procurement policy for the provision of medical equipment. o Include proper SLAs and NDAs to vendors and external associates | ID.GV-1 |  | ID.GV-1-02: le politiche di sicurezza IT di base sono definite, approvate dalla direzione e comunicate ai dipendenti e alle parti esterne pertinenti. es.: - Politica di utilizzo accettabile - Politica sulle password - Politica di gestione del cambiamento | ID.GV-1-02: Basic IT security policies are defined, approved by management, and communicated to employees and relevant external parties. ES.: - Acceptable Use Policy - Password policy - Change management policy
| Provide security awareness training for employees that operate critical devices and systems | PR.AT-1 |  | PR.AT-1-01: è fornita formazione sulla consapevolezza della sicurezza per i dipendenti che utilizzano dispositivi e sistemi critici | provided for employees using critical devices and systems
| and make background checks prior to authorizing access to key personnel. | PR.AC-6 |  | PR.AC-6-01: sono effettuati controlli in background prima di autorizzare l'accesso, su dispositivi critici, al personale chiave. | PR. AC-6-01: Background checks are carried out before authorizing access to key personnel on critical devices.
| Catalogue assets in an inventory of all medical devices, servers and workstations. - Assign unique ID’s to healthcare information systems, medical devices. | ID.AM-1 |  | ID.AM-1-01: sono catalogate in un inventario tutti i dispositivi medici, server e workstation. ID.AM-1-02: sono assegnati ID univoci a sistemi informativi sanitari, dispositivi medici. | ID.AM-1-01: All medical devices, servers and workstations are cataloged in an inventory. ID.AM-1-02: Unique IDs are assigned to health information systems, medical devices.
| Monitor and keep track of changes in ecosystem parties, so that business processes are not interrupted or hide risks. | PR.IP-3 |  | PR.IP-3-01: i cambiamenti nelle parti dell'ecosistema sono monitorati e ne è tenuta traccia, in modo che i processi aziendali non vengano interrotti o nascondano i rischi. | PR. IP-3-01: Changes in parts of the ecosystem are monitored and tracked, so business processes are not disrupted or risks are hidden.
| Apply the principle of least privilege to user workstations and connected devices. - Least privileges must also take into account data minimisation per role. | PR.AC-4 |  | PR.AC-4-01: è applicato il principio del privilegio minimo alle workstation degli utenti e ai dispositivi connessi. PR.AC-4-02: I privilegi minimi tengono conto della minimizzazione dei dati per ruolo. | PR. AC-4-01: The principle of least privilege applies to user workstations and connected devices. PR. AC-4-02: Least privileges take into account per-role data minimization.
| Data integrity should be ensured e.g. through hashing, integrity checks. | PR.DS-1 |  | PR.DS-1-03: L'integrità dei dati è garantita, ad es. tramite hashing, controlli di integrità. | PR. DS-1-03: Data integrity is ensured, e.g. through hashing, integrity checks.
| Establish appropriate security measures for the use of mobile devices and teleworking. | ID.GV-1 |  | ID.GV-1-03: sono stabilire adeguate misure di sicurezza per l'uso dei dispositivi mobili e il telelavoro. | ID.GV-1-03: Appropriate security measures are established for the use of mobile devices and teleworking.
| The Health Information System (HIS) must be able to monitor the correct operation of the equipment. - Monitor device behaviour in the context of medical workflows. | DE.CM-1 |  | DE.CM-1-01: Il Sistema Informativo Sanitario (HIS) è in grado di monitorare il corretto funzionamento delle apparecchiature. DE.CM-1-02: il comportamento del dispositivo nel contesto dei flussi di lavoro medici è monitorato. | DE. CM-1-01: The Health Information System (HIS) is able to monitor the correct functioning of the equipment. DE. CM-1-02: Device behavior in the context of medical workflows is monitored.
| Implement data recovery mechanisms to restore data from critical systems | RC.RP-1 |  | RC.RP-1-01: sono implementati meccanismi di recupero dati per ripristinarlii dai sistemi critici | RC. RP-1-01: Data recovery mechanisms are implemented to restore data from critical systems
| Investigate major incidents and review actions taken to mitigate and reduce time to react to future occurrences. | RS.IM-1 |  | RS.IM-1-01: gli incidenti gravi sono indagati RS.IM-1-02: sono riviste le azioni intraprese per mitigare gli incidenti gravi e ridurre i tempi di reazione a eventi futuri. | RS. IM-1-01: serious incidents are investigated RS. IM-1-02: Actions taken to mitigate serious accidents and reduce reaction times to future events are reviewed.
| Develop a disaster recovery plan, taking into account the minimum recovery requirements. | PR.IP-9 |  | PR.IP-9-01: è sviluppato un piano di ripristino di emergenza, tenendo conto dei requisiti minimi di ripristino. | PR. IP-9-01: A disaster recovery plan is developed, taking into account the minimum recovery requirements
| Avoid the use of End of life third-party components and devices on the operating environment, where possible take additional measures such as network isolation. | PR.AC-5; | da rivedere | PR.AC-5-06: componenti e dispositivi di terze parti fuori uso non sono usati nell'ambiente operativo PR.AC-5-07: adottare misure aggiuntive come l'isolamento della rete. | PR. AC-5-06: End-of-life third-party components and devices are not used in the operating environment PR. AC-5-07: Take additional measures such as network isolation.

## Annex I.II
Riferimento |  | Dettagli | Subcategory | Controlli | Controlli_EN
Part 1 – Governance and Ecosystem
Information System Security Governance & Risk Management | Information system security risk analysis | Conduct a risk assessment and impact assessment based on industry standards (like DICOM or ISO 80001 for medical devices or ISO 27799). Medical device introduction should be subject to such a risk assessment | ID.RA-1 | ID.RA-1-01: una valutazione del rischio e dell'impatto è condotta sulla base degli standard del settore (come DICOM o ISO 80001 per i dispositivi medici o ISO 27799). ID.RA-1-02: L'introduzione di dispositivi medici è soggetta a valutazione del rischio secondo il controllo ID.RA-1-01 | ID.RA-1-01: A risk and impact assessment is conducted based on industry standards (such as DICOM or ISO 80001 for medical devices or ISO 27799). ID.RA-1-02: The introduction of medical devices is subject to risk assessment according to control ID.RA-1-01
| Information system security policy | Adequate patch management processes - for medical devices - for the operating environment - Deployed in a timely manner | PR.IP-12 | PR.IP-12-01: sono eseguiti adeguati processi di gestione delle patch per i dispositivi medici PR.IP-12-02: sono eseguiti adeguati processi di gestione delle patch per l'ambiente operativo PR.IP-12-03: sono eseguiti e distruìibuiti adeguati processi di gestione delle patch in modo tempestivo | PR. IP-12-01: Appropriate patch management processes are in place for medical devices PR. IP-12-02: Appropriate patch management processes are in place for the operating environment PR. IP-12-03: Appropriate patch management processes are executed and destroyed in a timely manner
|  | Only use genuine software and ban all illegitimate software | PR.DS-6 | PR.DS-6-03: è utilizzato solo software originale e sono banditi tutti i software illegittimi | PR. DS-6-03: Only genuine software is used and all illegitimate software is banned
|  | Define a set of baseline IT security policies: - Management should approve the policies - Change Management policy - Establish a backup policy on selected critical systems. - Establish a vendor/procurement policy for the provision of medical equipment. - Include proper SLAs and NDAs to vendors and external associates | ID.GV-1 | ID.GV-1-04: La direzione approva le politiche ID.GV-1-05: è stabilita una Politica di gestione del cambiamento ID.GV-1-06: è stabilita una politica di backup su sistemi critici selezionati. ID.GV-1-07: è stabilita una politica di fornitore/appalto per la fornitura di attrezzature mediche. ID.GV-1-08: sono inclusi SLA e NDA adeguati a fornitori e collaboratori esterni | ID.GV-1-04: Management approves policies ID.GV-1-05: A Change Management Policy is established ID.GV-1-06: A backup policy is established on selected critical systems. ID.GV-1-07: A supplier/procurement policy for the supply of medical equipment is established. ID.GV-1-08: Includes SLAs and NDAs suitable for suppliers and contractors
|  | - Define clear roles for users of critical systems and medical devices, segregate duties | ID.AM-6 | ID.AM-6-02: sono definiti chiaramente, e separati, i ruoli per gli utenti di sistemi critici e dispositivi medici | ID.AM-6-02: Roles for users of critical systems and medical devices are clearly defined and separated.
|  | - Define information classification levels and label devices handling such information accordingly | ID.AM-5 | ID.AM-5-01: sono definiti i livelli di classificazione delle informazioni ID.AM-5-02: i dispositivi che gestiscono le informazioni del controllo ID.AM-5-01 sono etichettati di conseguenza | ID.AM-5-01: Information classification levels are defined ID.AM-5-02: Devices that handle ID.AM-5-01 control information are labeled accordingly
|  | policies communicated to employees and external parties | ID.GV-1 | ID.GV-1-09: le politiche sono comunicate ai dipendenti e ai soggetti esterni | ID.GV-1-09: Policies are communicated to employees and external parties
| Human resource security | Security awareness training for employees that operate critical devices and systems | PR.AT-1: | PR.AT-1-02: i dipendenti che utilizzano dispositivi e sistemi critici seguono corsi di formazione sulla sensibilizzazione alla sicurezza | PR. AT-1-02: Employees using critical devices and systems undergo security awareness training
|  | Background checks prior to authorizing access to key personnel | PR.AC-6 | PR.AC-6-02: prima di autorizzare l'accesso al personale chiave sono effettuati controlli in background | PR. AC-6-02: background checks are carried out before key personnel are granted access
| Asset Management | Catalogue medical devices and relevant assets in inventory. - Assign unique ID’s to healthcare information systems, medical devices | ID.AM-1 | ID.AM-1-03: sono catalogate in un inventario tutti i dispositivi medici, server e workstation. ID.AM-1-04: sono assegnati ID univoci a sistemi informativi sanitari, dispositivi medici. | ID.AM-1-03: All medical devices, servers and workstations are cataloged in an inventory. ID.AM-1-04: Unique IDs are assigned to health information systems, medical devices.
Part 2 – Protection
IT Security Architecture | Systems configuration | Operating environment must not hinder the application of security measures on the medical device or force the device to operate in lower security settings. | N/A
|  | Session management measures (e.g. session timeouts). | PR.AC-3 | PR.AC-3-01: sono implementate misure di gestione delle sessioni (ad es. timeout delle sessioni). | PR. AC-3-01: Session management measures (e.g. session timeouts) are implemented.
|  | Operating system hardening and application whitelisting | PR.DS-1, PR-IP-1 | PR.DS-1-04: è implementata la protezione avanzata del sistema operativo e la whitelist delle applicazioni | PR. DS-1-04: Hardened OS security and application whitelisting implemented
|  | Antivirus / anti-malware software | DE.CM-4 | DE.CM-4-01: è implementato un software antivirus/antimalware | DE. CM-4-01: Anti-virus/antimalware software is implemented
|  | Use of strong passwords | PR.DS-1: | PR.DS-1-05: sono utilizzate password complesse | PR. DS-1-05: Strong passwords are used
|  | Appropriate security measures for mobile devices and teleworking | N/A
| System segregation | Firewall | PR.AC-5: | PR.AC-5-08: è implementato un Firewall | PR. AC-5-08: a Firewall is implemented
|  | Network segmentation | PR.AC-5: | PR.AC-5-09: è implementata la Network segmentation | PR. AC-5-09: Network segmentation is implemented
|  | Partitioning mechanisms and traffic segmentation | PR.AC-5: | PR.AC-5-10: sono implementati meccanismi di partizionamento e segmentazione del traffico | PR. AC-5-10: traffic partitioning and segmentation mechanisms are implemented
| Traffic filtering | Use of traffic filtering software and hardware | PR.AC-5: | PR.AC-5-11: sono utilizzati software e hardware di filtraggio del traffico | PR. AC-5-11: traffic filtering software and hardware are used
| Cryptography | Encryption when storing sensitive personal data | PR.DS-1 | PR.DS-1-06: la crittografia è usata durante la memorizzazione di dati personali sensibili | PR. DS-1-06: Encryption is used when storing sensitive personal data
|  | Encryption of data in transit | PR.DS-2 | PR.DS-2-01: la crittografia è usata per i dati in transito | PR. DS-2-01: Encryption is used for data in transit
IT Security Administration | Administration information systems | Memory protection measures to block arbitrary code execution | PR.DS-1, PR-IP-1 | PR.DS-1-07: sono implementate misure di protezione della memoria per bloccare l'esecuzione di codice arbitrario | PR. DS-1-07: Memory protection measures are implemented to block arbitrary code execution
|  | Compatibility of medical device management software with security solutions that counter malicious code. | N/A
|  | Install only software programmes necessary for intended uses. | N/A
Identity and access management | Authentication and identification | User access management (credentials for accessing software applications or devices, user access policy etc.) | PR.AC-1 | PR.AC-1-03: è implementata la gestione degli accessi degli utenti (credenziali per l'accesso alle applicazioni o ai dispositivi software, policy di accesso degli utenti, ecc.) | PR. AC-1-03: user access management is implemented (credentials for access to applications or software devices, user access policies, etc.)
| Access rights | Apply principle of least privilege to user workstations and connected devices. | PR.AC-4 | PR.AC-4-03: è applicato il principio del privilegio minimo alle workstation degli utenti e ai dispositivi connessi. | PR. AC-4-03: Least-privilege applies to user workstations and connected devices.
|  | Least privileges must take into account data minimisation per role | PR.AC-4 | PR.AC-4-04: I privilegi minimi tengono conto della minimizzazione dei dati per ruolo | PR. AC-4-04: Least Privileges Accounts for Data Minimization by Role
IT security maintenance | IT security maintenance procedure | Provisions regarding patch management | PR.IP-7 | PR.AC-2-04: sono implementate disposizioni relative alla gestione delle patch | PR. AC-2-04: Patch management provisions are implemented
|  | Support patching without compromising interoperability/compatibility | PR.IP-7 | PR.AC-2-05: l'applicazione di patch è supportata senza compromettere l'interoperabilità/compatibilità | PR. AC-2-05: Patching is supported without compromising interoperability/compatibility
|  | Do not use End of life third-party components and devices on the operating environment | N/A
Physical and environmental security | Physical and environmental security | Regulated and authenticated physical access via suitable technical measures (e.g. badges) | PR.AC-2 | PR.AC-2-03: l'accesso fisico è regolamentato e autenticato tramite idonee misure tecniche (es. badge) | PR. AC-2-03: physical access is regulated and authenticated through appropriate technical measures (e.g. badges)
|  | Define roles and access rights, including those for physical access to medical devices | PR.AC-2 | PR.AC-2-04: sono definiti ruoli e diritti di accesso, compresi quelli per l'accesso fisico ai dispositivi medici | PR. AC-2-04: roles and access rights are defined, including those for physical access to medical devices
|  | Use of segregated, secure areas with appropriate access controls | N/A
Part 3 - Defence
Detection | Detection | Software integrity checks and device authentication mechanisms | PR.DS-6: | PR.DS-6-04: sono implementati controlli di integrità del software e meccanismi di autenticazione del dispositivo | PR. DS-6-04: Software integrity checks and device authentication mechanisms are implemented
|  | Data integrity ensured e.g. through hashing, integrity checks. | PR.DS-6: | PR.DS-6-05: l'integrità dei dati è garantita ad es. tramite hashing, controlli di integrità. | PR. DS-6-05: Data integrity is ensured e.g. through hashing, integrity checks.
| Logging | Monitor and keep track of changes in ecosystem parties, so that business processes are not interrupted or hide risks | PR.PT-1 | PR.PT-1-01: i cambiamenti nelle parti dell'ecosistema sono monitorati e ne è tenuta traccia, in modo che i processi aziendali non vengano interrotti o nascondano i rischi. | PR. PT-1-01: Changes in parts of the ecosystem are monitored and tracked, so business processes are not disrupted or hidden risks.
| Logs correlation and analysis | The Health Information System (HIS) must be able to monitor the correct operation of the equipment. - Monitor device behaviour in the context of medical workflows | DE.CM-1; PR.PT-1 | DE.CM-1-02: Il Sistema Informativo Sanitario (HIS) è in grado di monitorare il corretto funzionamento delle apparecchiature. DE.CM-1-03: il comportamento del dispositivo nel contesto dei flussi di lavoro medici è monitorato. PR.PT-1-01: E' definita una policy per la definizione del formato di log per le operazion di monitoraggio dell'HIS | DE. CM-1-02: The Health Information System (HIS) is able to monitor the correct functioning of the equipment. DE. CM-1-03: Device behavior in the context of medical workflows is monitored. PR. PT-1-01: A policy is defined for the definition of the log format for HIS monitoring operations
Part 4 - Resilience
Continuity of Operations | Business continuity management | Investigate major incidents and review actions taken to mitigate and reduce time to react to future occurrences | RS.IM-1 | RS.IM-1-03: gli incidenti gravi sono indagati RS.IM-1-04: sono riviste le azioni intraprese per mitigare gli incidenti gravi e ridurre i tempi di reazione a eventi futuri. | RS. IM-1-03: serious accidents are investigated RS. IM-1-04: Actions taken to mitigate serious incidents and reduce reaction times to future events are reviewed.
| Disaster recovery management | Develop a disaster recovery plan, taking into account the minimum recovery requirements | PR.IP-9 | PR.IP-9-02: è sviluppato un piano di ripristino di emergenza, tenendo conto dei requisiti minimi di ripristino. | PR. IP-9-02: A disaster recovery plan is developed, taking into account the minimum recovery requirements.
|  | Implement data recovery mechanisms to critical systems | RC.RP-1 | RC.RP-1-02: sono implementati meccanismi di recupero dati per ripristinarlii dai sistemi critici | RC. RP-1-02: Data recovery mechanisms are implemented to restore data from critical systems