## Foglio1
Riferimento |  | Dettagli | Subcategory | Controlli | Controlli_EN
Purchasing Controls | Security Audit/Evaluation | Prior to any medical device being purchased or brought onto any network the device should be compared to the organizations internal security standards and a determination should be made as to whether or not the device is capable of meeting those standards. | ID.RA-1; ID.AM-1 | ID.RA-1-01: Prima dell'acquisto o dell'immissione in rete di qualsiasi dispositivo medico, il dispositivo è confrontato con gli standard di sicurezza interni dell'organizzazione | ID.RA-1-01: Before purchasing or placing any medical device on the network, the device is compared to the organization's internal security standards
| Privacy Audit/Evaluation | In a similar manner, prior to any system acquisition, a privacy evaluation should be performed to ensure that the device possesses the requisite security controls to ensure that patient data is collected, stored, and transmitted in a way that is consistent with organizational policy. | ID.RA-1; ID.AM-1 | ID.RA-1-02: prima di qualsiasi acquisizione di un sistema, è eseguita una valutazione della privacy per garantire che il dispositivo disponga dei controlli di sicurezza necessari per garantire che i dati del paziente siano raccolti, archiviati e trasmessi in modo coerente con la politica dell'organizzazione. | ID.RA-1-02: Prior to any system acquisition, a privacy assessment is performed to ensure that the device has the necessary security controls in place to ensure that patient data is collected, stored, and transmitted in a manner consistent with the organization's policy.
| Support | Given the critical role that patching vulnerabilities plays in maintaining the security of any system, particular attention to should be given to what kind of support the vendor will provide for the software of devices, how frequently they release patches, and for how many years they will continue to provide patches for the device. Given the long active lifetime of many medical products, paying attention to the long-term ability to patch systems is a critical security consideration that should weigh in on purchasing decisions. Related to this consideration should be given to a vendors policies concerning security issues that are presented to them. | ID.SC-3 | ID.SC-3-01: prima di qualsiasi acquisizione di un sistema, è considerata la capacità a lungo termine di applicare patch ai sistemi ID.SC-3-02: sono fornite le politiche dei fornitori relative ai problemi di sicurezza che vengono loro presentati. | ID.SC-3-01: Before any acquisition of a system, the long-term ability to patch systems is considered ID.SC-3-02: Supplier policies regarding security issues presented to them are provided.
Perimeter Defenses | Firewalls | Firewalls at the perimeter are an essential control to ensure that communications between medical devices and external resources are either outright denied, where feasible, or restricted to just the communications that are essential for the device to function properly. | PR.AC-5 | PR.AC-5-01: sono implementati firewall lungo il perimetro per garantire che le comunicazioni tra dispositivi medici e risorse esterne siano completamente negate, ove possibile, o limitate alle sole comunicazioni essenziali per il corretto funzionamento del dispositivo. | PR. AC-5-01: Firewalls are implemented along the perimeter to ensure that communications between medical devices and external resources are completely denied, where possible, or limited to communications essential to the proper functioning of the device.
|  | In the case where a medical device is reachable over the internet, particular attention should be taken to ensure that the device has a separate administrative interface and that external access to the administrative interface is not possible from outside the organizations internal network. | PR.AC-5 | PR.AC-5-02: i dispositivi medici che sono raggiungibili tramite Internet, dispongono di un'interfaccia amministrativa separata e l'accesso all'interfaccia amministrativa non è possibile dall'esterno della rete interna dell'organizzazione. | PR. AC-5-02: Medical devices that are reachable via the Internet, have a separate administrative interface, and access to the administrative interface is not possible from outside the organization's internal network.
| Network Intrusion Detection System | Network Intrusion Detection Systems at the perimeter can be helpful in detecting exploit attempts coming from external parties as well as helpful in detecting traffic going to command and control sites and ransomware key generation sites. | DE.CM-1 | DE.CM-1-01: sono implementati sistemi di rilevamento delle intrusioni di rete nel perimetro che rilevano i tentativi di exploit provenienti da parti esterne. | DE. CM-1-01: Network intrusion detection systems are implemented at the perimeter that detect exploit attempts from external parties.
|  | As such Network intrusion detection systems at the perimeter can be useful for their potential for providing early warning of an attack attempt or successful compromise of a network enabled device. | DE.CM-1 | DE.CM-1-02: i sistemi di rilevamento delle intrusioni di rete nel perimetro forniscono un avviso precoce di un tentativo di attacco o di compromissione riuscita di un dispositivo abilitato alla rete. | DE. CM-1-02: Network intrusion detection systems at the perimeter provide early warning of a successful attack attempt or compromise of a network-enabled device.
| Proxy Server/Web Filter | For devices that communicate with external resources via http and/or https a proxy server or web filtering appliance may allow for even finer grained control over communications than a firewall. | PR.AC-5 | PR.AC-5-03: i dispositivi comunicano con risorse esterne tramite http e/o https, un server proxy o un'appliance di filtraggio Web consentendo un controllo ancora più dettagliato sulle comunicazioni rispetto a un firewall. | PR. AC-5-03: Devices communicate with external resources via http and/or https, a proxy server, or web filtering appliance, allowing even more fine-grained control over communications than a firewall.
|  | Where this is possible, it is recommended to use a different AV engine than the one used on internal endpoints as that will help to maximize the chance of successful detection for any malware vector. | PR.AC-5 | PR.AC-5-04: è utilizzato un motore AV diverso da quello utilizzato sugli endpoint interni, ciò consente di massimizzare le possibilità di rilevamento per qualsiasi vettore di malware. | PR. AC-5-04: A different AV engine is used than the one used on internal endpoints, which maximizes the chances of detection for any malware vector.
|  | Additionally, many appliances perform the ability to perform SSL stripping and these appliances can often be used as a part of a data loss prevention (DLP) system as a result. DLP may be advisable for use with medical devices that may require internet access in some form, but would not normally be used to transmit PII or PHI to an external entity. | PR.AC-5 | PR.AC-5-05: le appliance che eseguono lo stripping SSL sono utilizzate come parte di un sistema di prevenzione della perdita di dati (DLP). | PR. AC-5-05: Appliances that perform SSL stripping are used as part of a data loss prevention (DLP) system.
Network Security Controls | Network Segmentation | All medical devices should be an isolated network segment that restricts communication of the devices to just the systems that are required for the device to function. All other communications should be restricted. | PR.AC-5 | PR.AC-5-06: Tutti i dispositivi medici sono un segmento di rete isolato che limita la comunicazione dei dispositivi ai soli sistemi necessari per il funzionamento del dispositivo. PR.AC-5-07: Tutte le comunicazioni escluse dal controllo PR.AC-5-06 sono limitate. | PR. AC-5-06: All medical devices are an isolated network segment that limits device communication to only systems necessary for the operation of the device. PR. AC-5-07: All communications excluded from PR control. AC-5-06 are limited.
|  | Network segmentation is often achieved through the creation of VLANs and ACLs to control the flow of traffic in VLANs, but can also be achieved by using an separate physical network infrastructure, which is particularly useful in areas where a concentration of the same type of medical device will be deployed in a given area. | PR.AC-5 | PR.AC-5-08: La segmentazione della rete è ottenuta attraverso la creazione di VLAN e ACL per controllare il flusso di traffico nelle VLAN, PR.AC-5-09: La segmentazione della rete è ottenuta utilizzando un'infrastruttura di rete fisica separata | PR. AC-5-08: Network segmentation is achieved through the creation of VLANs and ACLs to control the flow of traffic in VLANs, PR. AC-5-09: Network segmentation is achieved using a separate physical network infrastructure
| Internal Firewalls | Internal firewalls can be used to improve upon network segmentation and used to further restrict communications of devices to just the systems (internal and external) that they need to interact with. | PR.PT-4; PR.AC-5 | PR.PT-4-01: sono utilizzati firewall interni per migliorare la segmentazione della rete e per limitare ulteriormente le comunicazioni dei dispositivi solo ai sistemi (interni ed esterni) con cui devono interagire. | PR. PT-4-01: Internal firewalls are used to improve network segmentation and to further restrict device communications to only the systems (internal and external) with which they need to interact.
|  | Firewalls, particularly next generation models, can also provide ways of monitoring and restricting traffic in ways that ACLs in switches cannot as they typically allow for deeper levels of traffic inspection. | PR.PT-4 | PR.PT-4-02: I firewall sono usati per monitorare e limitare il traffico in maniera più approfondita rispetti le ACL. | PR. PT-4-02: Firewalls are used to monitor and restrict traffic more thoroughly than ACLs.
|  | Internal firewalls are also highly useful for protecting “one-off” devices, such as an MRI machine, where isolation is sought but the presence of only a single device does not warrant the purchase of an entirely separate physical network infrastructure. | PR.PT-4; PR.AC-5 | PR.PT-4-03: i dispositivi "una tantum", come una macchina per la risonanza magnetica, sono protetti dai firewall interni | PR. PT-4-03: One-time devices, such as an MRI machine, are protected by internal firewalls
|  | Internal firewalls help to promote a zero trust model with regards to the communication with medical devices. | PR.PT-4; PR.AC-5 | PR.PT-4-04: è usato un modello zero trust per quanto riguarda la comunicazione con i dispositivi medici, grazie all'implementazione di firewall interni | PR. PT-4-04: Zero trust model is used for communication with medical devices, thanks to the implementation of internal firewalls
| Internal Network IDS | If traffic from network segments containing medical devices is routed through an internal network Intrusion Detection System, signatures can be created to detect default login credentials, attempts to connect to command and control IPs, and other forms of network traffic that may indicate an attack on a medical device or a successful compromise of a medical device. | DE.CM-1 | DE.CM-1-03: sono create firme per rilevare le credenziali di accesso predefinite, i tentativi di connessione agli IP di comando e controllo e altre forme di traffico di rete che indicano un attacco a un dispositivo medico o una compromissione riuscita di un dispositivo medico. | DE. CM-1-03: Signatures are created to detect default login credentials, attempts to connect to command and control IPs, and other forms of network traffic that indicate an attack on a medical device or a successful compromise of a medical device.
| Syslog Server | Where possible medical device logs should not just be stored on the device itself but should be exported to a distinct syslog server to allow for the collection and analysis of events that affect the device. | PR.IP-4 | PR.IP-4-01: i registri dei dispositivi medici sono archiviati anche su un server syslog distinto per consentire la raccolta e l'analisi degli eventi che interessano il dispositivo. | PR. IP-4-01: Medical device logs are also stored on a separate syslog server to allow collection and analysis of events affecting the device.
| Log Monitoring | Related to above control some form SIEM or log analysis should be performed on the collected log data. For example, a high occurrence of failed login attempts on a device or even a high occurrence of successful logins across a large number of devices (outside of scheduled maintenance) may be indicative of an attack from IoT malware like Mirai. | PR.PT-1 | PR.PT-1-01: In relazione al controllo PR.IP-4-01 sui dati di registro raccolti sono eseguiti alcuni moduli SIEM o analisi del registro. Ad esempio, un numero elevato di tentativi di accesso non riusciti su un dispositivo o anche un numero elevato di accessi riusciti su un numero elevato di dispositivi (al di fuori della manutenzione programmata) può essere indicativo di un attacco da malware IoT come Mirai. | PR. PT-1-01: In relation to PR control. IP-4-01 on the collected log data are performed some SIEM modules or log analysis. For example, a high number of failed login attempts on a device or even a large number of successful logins on a large number of devices (outside of scheduled maintenance) can be indicative of an IoT malware attack such as Mirai.
| Vulnerability Scanning | IoT devices should be routinely scanned to ensure that they are properly configured and that out of date software does not leave them susceptible to compromise. | DE.CM-8 | DE.CM-8-01: I dispositivi IoT sono scansionati di routine per garantire che siano configurati correttamente e che il software non aggiornato non li renda suscettibili di compromissione. | DE. CM-8-01: IoT devices are routinely scanned to ensure that they are configured correctly and that outdated software does not make them susceptible to compromise.
|  | As such IoT devices should be included in part of the larger vulnerability management program that the organization has in place. | PR.IP-12 | PR.IP-12-01: i dispositivi IoT sono inclusi inel programma di gestione delle vulnerabilità che l'organizzazione ha messo in atto. | PR. IP-12-01: IoT devices are included in the vulnerability management program that the organization has put in place.
|  | Not all IoT devices and features may be readily assessed by traditional vulnerability scanners and specialized scanners my have to be considered. | DE.CM-8 | DE.CM-8-02: i dispositivi IoT che non possono essere valutati da scanner di vulnerabilità tradizionali, sono valutati da scanner specializzati | DE. CM-8-02: IoT devices that cannot be evaluated by traditional vulnerability scanners, are evaluated by specialized scanners
|  | Even with specialized scanners there is such diversity amongst IoT devices that manual compliance auditing may be needed in some cases as indicated in the device security section below. | DE.CM-8 | DE.CM-8-03: i dispositivi IoT che non possono essere valutati da scanner specializzati, sono controllati manualmente, come indicato nella sezione sulla sicurezza del dispositivo. | DE. CM-8-03: IoT devices that cannot be evaluated by specialized scanners, are manually checked, as indicated in the device security section.
| DNS Sinkholes | The security of a medical device deployment can be improved by having dedicated DNS servers for the device that can only resolve the limited number of IP addresses required for the device to function. All other DNS requests can be sinkholed. | PR.PT-4 | PR.PT-4-05: i dispositivi medici dispongono di server DNS dedicati per in grado di risolvere solo il numero limitato di indirizzi IP richiesti per il funzionamento del dispositivo. Tutte le altre richieste DNS possono essere sprofondate. | PR. PT-4-05: Medical devices have dedicated DNS servers that can resolve only the limited number of IP addresses required for the device to operate. All other DNS requests can be sunk.
Device Security Controls | Change Default Credentials | All devices should have their default credentials changed prior to deployment on the network and devices with hardcoded credentials should not be used. | PR.AC-1 | PR.AC-1-01: Tutti i dispositivi hanno le credenziali predefinite modificate prima della distribuzione sulla rete PR.AC-1-02: i dispositivi con credenziali hardcoded non sono utilizzati. | PR. AC-1-01: All devices have default credentials changed before deployment on the network PR. AC-1-02: Devices with hard-coded credentials are not used.
|  | Account credentials used in place of the defaults should be compliant with organizational password polices. | PR.AC-1 | PR.AC-1-03: Le credenziali dell'account utilizzate al posto delle impostazioni predefinite sono conformi ai criteri delle password dell'organizzazione. | PR. AC-1-03: Account credentials used instead of default settings comply with your organization's password policy.
| Account Lockout | Account lockout features should be configured to block logins after 3-5 login attempts. | PR.AC-1 | PR.AC-1-04: Le funzionalità di blocco dell'account sono configurate per bloccare gli accessi dopo 3-5 tentativi di accesso. | PR. AC-1-04: Account lockout features are configured to block access after 3-5 login attempts.
| Enable Secure Transport | Devices should be configured to only send data in a secure format and secure protocols like ssh and https should be used in place of insecure protocols like telnet and http. | PR.DS-2 | PR.DS-2-01: I dispositivi sono configurati per inviare dati solo in un formato sicuro PR.DS-2-02: sono usati esclusivamnete protocolli sicuri come ssh e https al posto di protocolli non sicuri come telnet e http. | PR. DS-2-01: Devices are configured to send data only in a secure format PR. DS-2-02: Secure protocols such as ssh and https are used exclusively in place of insecure protocols such as telnet and http.
|  | Insecure networking protocols should be disabled wherever possible. | PR.DS-2 | PR.DS-2-03: I protocolli di rete non sicuri sono disabilitati ove possibile. | PR. DS-2-03: Insecure network protocols are disabled where possible.
| Spare copy of firmware/software | In the event that a device is compromised or runs into some other software issue, having a spare copy of the devices firmware or software is critical to restoring the device to functional state in a timely manner. | PR.IP-4 | PR.IP-4-02: Nel caso in cui un dispositivo sia compromesso o si imbatta in qualche altro problema software, si dispone di una copia di riserva del firmware o del software del dispositivo. | PR. IP-4-02: In case a device is compromised or runs into some other software problem, you have a backup copy of the device's firmware or software.
|  | Staff should be trained and competent in procedures to load reload software/firmware on the various kinds of medical devices supported. | PR.AT-1 | PR.AT-1-01: Il personale è formato e competente nelle procedure per caricare il software/firmware di ricarica sui vari tipi di dispositivi medici supportati. | PR. AT-1-01: Staff are trained and competent in the procedures for uploading charging software/firmware to the various types of supported medical devices.
| Backup of device configuration | Backing up these custom settings after changes occur will help to ensure that devices can be restored to functional status in as timely a manner as possible. | PR.IP-4 | PR.IP-4-03: vengono eseguiti backup delle impostazioni personalizzate per ripristinare il dispositivo medico allo stato funzionale nel modo più tempestivo possibile. | PR. IP-4-03: Custom settings are backed up to restore the medical device to functional status in the most timely manner possible.
| Baseline Configurations | Related to the controls above baseline configurations should be established for each device that ensure the proper configuration of the device with regards to clinical functionality and security. | PR.IP-1 | PR.IP-1-01: sono stabilite configurazioni di base per ciascun dispositivo medico | PR. IP-1-01: basic configurations are established for each medical device
|  | In the event a device specific backup is not available this baseline configuration can be modified to ensure the quick restoration of device in a manner that is compliant with organizational security policies. | PR.IP-1 | PR.IP-1-02: la configurazione di base di un dispositivo medico è modificata per garantirne il ripristino rapido in modo conforme alle politiche di sicurezza dell'organizzazione, solo nel caso in cui non sia disponibile un backup specifico del dispositivo, | PR. IP-1-02: The basic configuration of a medical device is modified to ensure rapid recovery in accordance with the organization's security policies, only if a device-specific backup is not available
| Encrypt Storage | Medical devices should support encryption of any PHI and/or PII stored on the device. | PR.DS-1 | PR.DS-1-01: I dispositivi medici supportano la crittografia di qualsiasi PHI e/o PII memorizzata sul dispositivo. | PR. DS-1-01: Medical devices support encryption of any PHI and/or PII stored on the device.
|  | This feature should be turned on in case of device theft or an unauthorized user gaining physical access to the device. | PR.DS-1 | PR.DS-1-02: Questa funzione deve essere attivata in caso di furto del dispositivo o in caso di accesso fisico al dispositivo da parte di un utente non autorizzato. | PR. DS-1-02: This feature must be enabled if the device is stolen or if an unauthorized user has physically accessed the device.
| Different User Accounts | Admin accounts and user level accounts should be possible and ideally the admin account should be bound to the management interface and unusable on any internet facing interface. | PR.AC-4 | PR.AC-4-01: Gli account amministratore e gli account a livello di utente sono possibili PR.AC-4-02: l'account amministratore è collegato all'interfaccia di gestione ed è inutilizzabile su qualsiasi altra interfaccia con connessione Internet. | PR. AC-4-01: Administrator Accounts and User-Level Accounts Are Possible PR. AC-4-02: The administrator account is linked to the management interface and is unusable on any other interface with an Internet connection.
| Restrict Access to Management Interface | The management interface of the device has the potential to do the most damage to the device if compromised, as it will more easily allow access to the administrative functions of the device. Communication to this interface should be locked down to only authorized terminals for making changes to the device. | PR.AC-4 | PR.AC-4-03: La comunicazione con l'interfaccia di gestione è possibile solo ai terminali autorizzati per apportare modifiche al dispositivo. | PR. AC-4-03: Communication with the management interface is only possible for authorized terminals to make changes to the device.
| Update Mechanisms | Mechanisms should be put in place to identify the need for updates and to ensure the routine update of all medical devices so that unpatched vulnerabilities remain minimized. | PR.IP-12 | PR.IP-12-02: sono messi in atto meccanismi per identificare la necessità di aggiornamenti e per garantire l'aggiornamento di routine di tutti i dispositivi medici. | PR. IP-12-02: Mechanisms are put in place to identify the need for updates and to ensure routine updating of all medical devices.
| Compliance Monitoring | Compliance monitoring should be routinely done to ensure that updates or other changes keep device configurations consistent with baseline configurations and organizational security policies. | PR.IP-3 | PR.IP-3-01: Il monitoraggio della conformità è eseguito regolarmente per garantire che gli aggiornamenti o altre modifiche mantengano le configurazioni dei dispositivi coerenti con le configurazioni di base e le politiche di sicurezza dell'organizzazione. | PR. IP-3-01: Compliance monitoring is performed regularly to ensure that updates or other changes keep device configurations consistent with your organization's basic configurations and security policies.
| Physical Security | Security controls should be put in place to ensure that physical access to medical devices is limited only to authorized individuals and that physical theft of the device is prohibited. | PR.AC-2 | PR.AC-2-01: sono messi in atto controlli di sicurezza per garantire che l'accesso fisico ai dispositivi medici sia limitato solo alle persone autorizzate e che sia vietato il furto fisico del dispositivo. | PR. AC-2-01: Security controls are put in place to ensure that physical access to medical devices is restricted to authorized persons only and physical theft of the device is prohibited.
| Asset Management | Keeping track of which devices are in which locations and what versions of software/firmware they are all running will be invaluable in helping to determine the scope of potential incidents. | ID.AM-1; ID.AM-2 | ID.AM-1-01: è tenuta traccia di quali dispositivi si trovano in quali posizioni e quali versioni di software/firmware sono in esecuzione. | ID.AM-1-01: Track which devices are in which locations and which software/firmware versions are running.
Interface and Central Station Security | OS Hardening | Since this guide is specific to providing guidance on medical devices it will not go into depth on OS Hardening techniques, but techniques like the removal of unnecessary services, password protection, installation of AV, and other common OS hardening techniques should all be employed. | PR.DS-1, PR-IP-1 | PR.DS-1-01: sono utilizzate tecniche come la rimozione di servizi non necessari, la protezione con password, l'installazione di AV e altre tecniche comuni di protezione avanzata del sistema operativo. | PR. DS-1-01: Techniques such as removing unnecessary services, password protection, installing AV, and other common OS hardening techniques are used.
| Encrypted Transport | As with the medical devices themselves, these systems will be used to send and receive data and as such should make use of the same secure protocols discussed in the device configuration section. | PR.DS-2 | PR.DS-2-04: i trasporti crittografati utilizzati per inviare e ricevere dati e utilizzano gli stessi protocolli di sicurezza dispositivi medici. | PR. DS-2-04: Encrypted transports used to send and receive data and use the same medical device security protocols.
| Message Security – HL7 v3 Security Standards | The exchange of HL7 messages should be done using the HL7 v3 standard as this provides for security provisions not present in earlier versions of the HL7 standard. | PR.DS-2 | PR.DS-2-04: Lo scambio di messaggi HL7 è effettuato utilizzando lo standard HL7 v3. | PR. DS-2-04: HL7 message exchange is carried out using the HL7 v3 standard.
Security Testing | Penetration Testing | A penetration test can be an effective mean of assessing how effective your device and network configurations are at turning back an attack on medical device installed on your network. | PR.IP-2 | PR.IP-2-01: l'efficacia del dispositivo e delle configurazioni di rete nel respingere un attacco è valutata con un penetration test. | PR. IP-2-01: The effectiveness of the device and network configurations in repelling an attack is assessed with a penetration test.
|  | The results can be used to help you further improve your defenses and may reveal flaws in the device that can be presented to the manufacturer for patching in a upcoming update release. | PR.IP-7 | PR.IP-7-01: I risultati del penetration test sono utilizzati per migliorare ulteriormente le tue difese PR.IP-7-02:I risultati del penetration test sono utilizzati per rivelare difetti nel dispositivo. | PR. IP-7-01: Penetration test results are used to further improve your defenses PR. IP-7-02:I Penetration test results are used to reveal defects in the device.
Incident Response | Incident Response Plan | Organizations should have detailed plans in place to deal with the compromise of medical devices before such an incident becomes a reality for the organization. | PR.IP-9 | PR.IP-9-01: Le organizzazioni dispone di piani dettagliati per affrontare la compromissione dei dispositivi medici prima che un incidente diventi una realtà per l'organizzazione. | PR. IP-9-01: Organizations have detailed plans to address the compromise of medical devices before an incident becomes a reality for the organization.
|  | Organizations should have a clear cut plan in place that defines how they will react to an incident and who will be responsible for what actions during the detection, containment, eradication, and recovery phases. | PR.IP-9: | PR.IP-9-02: Le organizzazioni dispongono di un piano ben definito che definisce come reagire ad un incidente e chi sarà responsabile di quali azioni durante le fasi di rilevamento, contenimento, eradicazione e recupero. | PR. IP-9-02: Organizations have a well-defined plan that defines how to react to an incident and who will be responsible for what actions during the detection, containment, eradication and recovery phases.
|  | It is also important all staff are made aware of the plan and are trained to respond appropriately and effectively. | PR.AT-1 | PR.AT-1-02: tutto il personale sia informato del piano di risposta è formato per rispondere in modo appropriato ed efficace. | PR. AT-1-02: All staff are informed of the response plan shall be trained to respond appropriately and effectively.
| Mock Incidents | It would be highly beneficial for any organization to conduct a mock incident regarding the compromise of medical devices to ensure that they have an effective incident response plan in place and that employees are adept at carrying out that incident response plan. Mock incidents provide a great way to identify security deficiencies as well as effective practices and use the lessons learned to further improve your organization’s security posture. | RS.RP-1 | RS.RP-1-01: è condotto un finto incidente riguardante la compromissione di dispositivi medici per garantire che dispongano di un piano di risposta agli incidenti efficace e che i dipendenti siano abili nell'esecuzione di tale piano di risposta agli incidenti. | RS. RP-1-01: A mock incident involving the compromise of medical devices is conducted to ensure that they have an effective incident response plan in place and that employees are skilled in executing that incident response plan.